## RecallTrackerReport
Recall Tracking Report
Begin Date: 7/11/2025
End Date: 8/11/2025
Create Date | Complete Date | Facility | Resources | Priority | Recall Type | Recall Label | SmartMatch | Quantity in Hospital | Comments | Status
2025-07-16 13:42 | 2025-07-17 06:01 | New York Eye and Ear Infirmary (NYEEI) | Jovel Calderon Jannell Maldonado |  | Medical Device | ABBOTT - i-STAT TriControls Calibration Verification Set – IMPORTANT INFORMATION | No |  | No affected lot # on hand - 7/17/2025 Jovel Calderon Completed by: Jovel Calderon | Completed
|  | Mount Sinai Hospital (MSH/East) |  |  | Medical Device | Flexicare Medical (Dongguan) Ltd. - BritePro Solo SingeUse Fiber Optic Handle and Blade - Class 1 Recall | Yes |  | Duplicate from July 2025 This recall was removed by Christopher Paradise on 8/4/2025 | Deleted
|  | Mount Sinai Queens (MSQ) |  |  | Medical Device | Siemens Healthcare Diagnostics Inc - RAPIDPoint 500 Systems Wash/Waste Cartridge - Class 2 Recall | Yes |  | No affected items found in POH This recall was removed by Christopher Paradise on 8/4/2025 | Deleted
2025-07-17 07:47 | 2025-07-18 09:57 | South Nassau Communities Hospital (MSSN) | Michael Arabia Camille Cox Anthony Battaglia Eduardo Corona |  | Medical Device | Flexicare Medical Limited - URGENT: VOLUNTARY MEDICAL DEVICE RECALL BritePro Solo Fiber Optic Laryngoscope Handles | No |  | Not an inventory item - 7/18/2025 Anthony Battaglia Completed by: Christopher Paradise | Completed
|  | Mount Sinai Hospital (MSH/East) |  |  | Medical Device | Boston Scientific Corporation - Torflex Transseptal Guiding Sheath Kit - Class 2 Recall | Yes |  | Received on 6/17. Completed and closed. This recall was removed by Christopher Paradise on 7/14/2025 | Deleted
2025-07-16 13:36 | 2025-07-16 13:37 | Mount Sinai Hospital (MSH/East) | Rebeca Melendez Christopher Paradise |  | Medical Device | GE - Wall Stand Cable Preventive Maintenance for Certain Proteus XR-a Systems – CORRECTION | No |  | No affected product found in POH. Sent to Biomed for awareness. - 7/16/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
2025-08-04 09:25 | 2025-08-04 09:25 | Mount Sinai Hospital (MSH/East) | Rebeca Melendez Christopher Paradise |  | Medical Device | Philips Respironics - Continuous Ventilator Correction: Philips Respironics Updates Use Instructions for BiPAP A30, A40, and V30 Devices Due to Interruptions and/or Loss of Therapy | No |  | No items found in POH - 8/4/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
|  | Mount Sinai Hospital (MSH/East) |  |  | Medical Device | Olympus Corporation of the Americas - Glass Body For 194 Ellik Evacuator - Class 2 Recall | Yes |  | No affected product found on POH This recall was removed by Christopher Paradise on 7/23/2025 | Deleted
|  | Mount Sinai Morningside (MSM) |  |  | Medical Device | Edwards Lifesciences, LLC - Optisite Arterial Cannula - Class 1 Recall | Yes |  | Received on 5/23. Completed and closed. This recall was removed by Christopher Paradise on 7/14/2025 | Deleted
2025-07-23 12:02 | 2025-07-28 15:55 | Mount Sinai Hospital (MSH/East) | Juanita Jordan Jesus Mercado Pedro Rosario Leandro Romero Arbinda Gurung |  | Medical Device | BD URGENT Medical Device Recall SureStep™ Foley Tray Systems: Bardex® I.C. Complete Care® Infection Control Foley Catheter Tray; Lubri-Sil® I.C. Complete Care® Infection Control Foley Catheter Tray; BARD® Lubri-Sil® Foley Catheter Tray | No |  | None of the lots were affected - 7/23/2025 Arbinda Gurung Completed by: Pedro Rosario | Completed
2025-07-22 08:30 | 2025-07-24 06:06 | Mount Sinai-Union Square (MSUS/Downtown-Formly PACC) | Jesse Carbo Jovel Calderon Jannell Maldonado |  | Medical Device | CooperSurgical, Inc. - Endosee System Convenience Kit with IV Tube - Class 2 Recall | No |  | no affected item on hand - 7/24/2025 Jovel Calderon Completed by: Jovel Calderon | Completed
|  | Mount Sinai Queens (MSQ) |  |  | Medical Device | Volcano Corp - Visions PV .014P RX Digital IVUS Catheter, Visions PV .018 Digital IVUS Catheter - Class 2 Recall | Yes |  | Duplicate from 6/25 This recall was removed by Christopher Paradise on 8/1/2025 | Deleted
|  | Mount Sinai Hospital (MSH/East) |  |  | Medical Device | Boston Scientific Corporation - VersaCross Access Solution - Class 2 Recall | Yes |  | Received on 6/17. Completed and closed with no affected items found This recall was removed by Christopher Paradise on 7/14/2025 | Deleted
|  | Mount Sinai Morningside (MSM) |  |  | Medical Device | Edwards Lifesciences, LLC - FEMFLEX II - Class 1 Recall | Yes |  | Received on 5/23. Completed and closed. This recall was removed by Christopher Paradise on 7/14/2025 | Deleted
2025-07-23 08:55 | 2025-07-29 09:13 | Mount Sinai Hospital (MSH/East) | Maura Carpo Jason Smith Ishmael Cobbold |  | Medical Device | Olympus Corporation of the Americas - Blue Silicone - Class 2 Recall | No |  | Recall notice was sent to SPD management to address. - 7/29/2025 Jason Smith Completed by: Jason Smith | Completed
|  | Mount Sinai West (MSW) |  |  | Medical Device | Stryker Corporation - Stryker SmartPump Tourniquet - Class 2 Recall | Yes |  | DUPLICATE : NO AFFECTED PRODUCTS IN POH COMPLETED ON 7/3 This recall was removed by Rebeca Melendez on 8/8/2025 | Deleted
|  | Mount Sinai Morningside (MSM) |  |  | Medical Device | Edwards Lifesciences, LLC - FEMFLEX II - Class 1 Recall | Yes |  | Received on 5/23. Completed and closed. This recall was removed by Christopher Paradise on 7/14/2025 | Deleted
|  | New York Eye and Ear Infirmary (NYEEI) |  |  | Medical Device | Flexicare Medical (Dongguan) Ltd. - BritePro Solo Mini SingleUse Fiber Optic Laryngoscope Handle - Class 1 Recall | Yes |  | no affected lot# on inventory This recall was removed by Jovel Calderon on 8/4/2025 | Deleted
2025-08-04 05:55 |  | Mount Sinai Brooklyn Hospital (MSB) | Russell Smith - GE Ronald Narchet Russell Smith - MS |  | Medical Device | Siemens Healthcare Diagnostics Inc - RAPIDPoint 500 Systems Wash/Waste Cartridge - Class 2 Recall | Yes |  | We could have affected items from our purchase orders dating back to 2023. Please check to see if we have any of the affected items, and if so, how many? - 8/4/2025 Christopher Paradise We could have affected items from our purchase orders dating back to 2023. Please check to see if we have any of the affected items, and if so, how many? - 8/4/2025 Christopher Paradise | In Progress
|  | Icahn School of Medicine (ISOM) |  |  | Medical Device | B BRAUN MEDICAL INC - Extension Set - Class 1 Recall | Yes |  | DUPLICATE This recall was removed by Christopher Paradise on 8/8/2025 | Deleted
|  | Mount Sinai Brooklyn Hospital (MSB) |  |  | Medical Device | Ortho-Clinical Diagnostics, Inc. - VITROS Chemistry Products Performance Verifier I - Class 2 Recall | Yes |  | No affected items found in POH. Shared with Lab for awareness. This recall was removed by Christopher Paradise on 7/18/2025 | Deleted
|  | South Nassau Communities Hospital (MSSN) |  |  | Medical Device | Flexicare Medical (Dongguan) Ltd. - BritePro Solo SingeUse Fiber Optic Handle and Blade - Class 1 Recall | Yes |  | Duplicate from July 2025 This recall was removed by Christopher Paradise on 8/4/2025 | Deleted
2025-08-08 09:28 | 2025-08-08 09:29 | Mount Sinai Hospital (MSH/East) | Rebeca Melendez Christopher Paradise |  | Medical Device | AmeriSource Bergen / BBraun Recall Lactated Ringers Bag & Sodium Chloride - Due to the potential for fluid leakage at one of the weld sites. | No |  | No affected items found in POH - 8/8/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
|  | Mount Sinai Morningside (MSM) |  |  | Medical Device | Maquet Cardiopulmonary Gmbh - Infusion Pole - Class 2 Recall | Yes |  | No affected items found in POH This recall was removed by Christopher Paradise on 7/30/2025 | Deleted
2025-07-15 13:16 | 2025-07-15 13:17 | Mount Sinai Hospital (MSH/East) | Rebeca Melendez Christopher Paradise |  | Medical Device | BD URGENT MEDICAL DEVICE PRODUCT RECALL (MMS-25-5283-FA) FR-110 BEZEL ASSEMBLY FAILURE Type of Action: Removal | No |  | No affected product found in POH - 7/15/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
|  | Mount Sinai West (MSW) |  |  | Medical Device | Olympus Corporation of the Americas - Glass Body For 194 Ellik Evacuator - Class 2 Recall | Yes |  | No affected product found on POH This recall was removed by Christopher Paradise on 7/23/2025 | Deleted
2025-07-16 13:42 | 2025-07-23 12:52 | Mount Sinai Hospital (MSH/East) | Marie Grace Meghal Desai Radhika Banu |  | Medical Device | ABBOTT - i-STAT TriControls Calibration Verification Set – IMPORTANT INFORMATION | No |  | Please see response from the POCT team, attached. This recall does not impact MSH as ionized calcium values are not reported. - 7/23/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
2025-07-14 09:17 | 2025-07-14 11:05 | Mount Sinai Queens (MSQ) | Vanessa Palma-Tejeda Claire Jazmin-Aquino Victor Mena Eleana Candelier |  | Medical Device | Getinge - URGENT MEDICAL DEVICE – Correction - Atrium Ocean, Oasis, and Express Chest Drains | No | 29 | Signs are currently being posted in all units for this device correction. - 7/14/2025 Vanessa Palma-Tejeda Completed by: Vanessa Palma-Tejeda | Completed
2025-07-17 07:47 | 2025-07-22 12:23 | Mount Sinai Hospital (MSH/East) | Joshua Villar Michael Arabia Jesus Mercado Pedro Rosario |  | Medical Device | Flexicare Medical Limited - URGENT: VOLUNTARY MEDICAL DEVICE RECALL BritePro Solo Fiber Optic Laryngoscope Handles | No |  | No product on hand - 7/22/2025 Michael Arabia Completed by: Pedro Rosario | Completed
|  | Beth Israel Medical Center (BIMC) |  |  | Medical Device | Ortho-Clinical Diagnostics, Inc. - VITROS Chemistry Products Performance Verifier II - Class 2 Recall | Yes |  | Hospital in now closed. This recall was removed by Jovel Calderon on 7/18/2025 | Deleted
2025-07-16 07:53 | 2025-07-18 11:12 | Icahn School of Medicine (ISOM) | Raquel Padilla Torrian Lee Christian Chino |  | Medical Device | BD - URGENT MEDICAL DEVICE RECALL (CORRECTION) (MMS-25-5311-FA) Revised performance data for BD Alaris™ Pump Module Model 8100 with Guardrails™ Suite MX software when used with a subset of BD Alaris™ Compatible Pump Infusion Sets (Attachment A) Type of Action: Medical Device Correction | No |  | Icahn School of Medicine (ISOM) have been identified as affected sites in the BD Urgent Medical Device Recall (Correction), Reference Number MMS-25-5311-FA. The facilities have been notified accordingly. This communication serves to ensure awareness across impacted locations and promote timely coordination in response to the recall. - Russell Smith - 7/18/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
|  | Mount Sinai Queens (MSQ) |  |  | Medical Device | Stryker Corporation - Strykeflow - Class 2 Recall | Yes |  | Duplicate from 7/2/25 This recall was removed by Christopher Paradise on 7/25/2025 | Deleted
2025-07-23 12:02 | 2025-07-25 09:39 | Mount Sinai Morningside (MSM) | Adelsi Soto Jazzmin Gamble Erika Saldana |  | Medical Device | BD URGENT Medical Device Recall SureStep™ Foley Tray Systems: Bardex® I.C. Complete Care® Infection Control Foley Catheter Tray; Lubri-Sil® I.C. Complete Care® Infection Control Foley Catheter Tray; BARD® Lubri-Sil® Foley Catheter Tray | No |  | No affected products found. - 7/25/2025 Jazzmin Gamble Completed by: Erika Saldana | Completed
|  | Mount Sinai West (MSW) |  |  | Medical Device | Flexicare Medical (Dongguan) Ltd. - BritePro Solo SingeUse Fiber Optic Handle and Blade - Class 1 Recall | Yes |  | Duplicate from July 2025 This recall was removed by Christopher Paradise on 8/4/2025 | Deleted
|  | Mount Sinai-Union Square (MSUS/Downtown-Formly PACC) |  |  | Medical Device | Stryker Corporation - Stryker SmartPump Tourniquet - Class 2 Recall | Yes |  | No affected inventory on hand This recall was removed by Jovel Calderon on 8/8/2025 | Deleted
|  | Mount Sinai Hospital (MSH/East) |  |  | Medical Device | Olympus Corporation of the Americas - Ellik Evacuator Adapter, Locking - Class 2 Recall | Yes |  | No affected items found in POH This recall was removed by Christopher Paradise on 7/23/2025 | Deleted
2025-07-14 09:17 | 2025-07-15 08:22 | South Nassau Communities Hospital (MSSN) | Camille Cox Anthony Battaglia Eduardo Corona |  | Medical Device | Getinge - URGENT MEDICAL DEVICE – Correction - Atrium Ocean, Oasis, and Express Chest Drains | No |  | Received updated IFU - 7/15/2025 Anthony Battaglia Completed by: Christopher Paradise | Completed
2025-08-06 12:06 | 2025-08-06 12:31 | South Nassau Communities Hospital (MSSN) | Rasdeep Sappal |  | Medical Device | Cook Medical - MEDICAL DEVICE RECALL EFF,􀀁TIVENESS CHECK URGENT: PROMPT RESPO􀀁SE REQUIRED - Cook Beacon® Tip 5.0 Fr Angi1ographic Catheters | No |  | As of right now, we do not have any of the affected LOT numbers. Just looked at all our inventory in stock and those LOT numbers do not exist. - 8/6/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
|  | Mount Sinai Morningside (MSM) |  |  | Medical Device | Ortho-Clinical Diagnostics, Inc. - VITROS Chemistry Products Performance Verifier II - Class 2 Recall | Yes |  | No affected items found in POH. Shared with Lab for awareness. This recall was removed by Christopher Paradise on 7/18/2025 | Deleted
2025-08-04 05:59 |  | Icahn School of Medicine (ISOM) | Christopher Chan Julie Sazant |  | Medical Device | Remel, Inc - Blood Agar,5 Sheep Blood - Class 2 Recall | No |  | 1 PO found in 2025 at ISOM with possible affected product. Manufacturer Cat Number: R01202 - 8/4/2025 Christopher Paradise | In Progress
2025-07-22 08:19 | 2025-07-29 10:25 | Mount Sinai Hospital (MSH/East) | Maura Carpo Jason Smith Ishmael Cobbold |  | Medical Device | AbbVie - URGENT MEDICAL DEVICE RECALL - UPDATE Keller FunnelTM 2 | No |  | This recall was circulated back in March 2024. They have now added a lot number to the affected lots. Please check for lot number: 24C12C - 7/22/2025 Christopher Paradise No affected lot #'s in inventory. - 7/29/2025 Jason Smith Completed by: Jason Smith | Completed
|  | Mount Sinai Hospital (MSH/East) |  |  | Medical Device | Volcano Corp - Visions PV .014P RX Digital IVUS Catheter, Visions PV .018 Digital IVUS Catheter - Class 2 Recall | Yes |  | Duplicate from 6/25 This recall was removed by Christopher Paradise on 8/1/2025 | Deleted
|  | Mount Sinai Morningside (MSM) |  |  | Medical Device | Siemens Healthcare Diagnostics Inc - RAPIDPoint 500 Systems Wash/Waste Cartridge - Class 2 Recall | Yes |  | No affected items found in POH This recall was removed by Christopher Paradise on 8/4/2025 | Deleted
|  | Mount Sinai Queens (MSQ) |  |  | Medical Device | Olympus Corporation of the Americas - Blue Silicone - Class 2 Recall | Yes |  | Duplicate This recall was removed by Christopher Paradise on 7/23/2025 | Deleted
|  | Mount Sinai Queens (MSQ) |  |  | Medical Device | Flexicare Medical (Dongguan) Ltd. - BritePro Solo SingeUse Fiber Optic Handle and Blade - Class 1 Recall | Yes |  | Duplicate from July 2025 This recall was removed by Christopher Paradise on 8/4/2025 | Deleted
2025-07-16 12:43 | 2025-07-16 12:44 | Mount Sinai Hospital (MSH/East) | Rebeca Melendez Christopher Paradise |  | Medical Device | Medtronic - Urgent Medical Device Removal Update McGRATH™ MAC Video Laryngoscope Item Codes 300-000-000, 300-200-000 Devices Beyond the 5-Year Useful Life | No |  | No affected items found in POH. Went back 6 years. - 7/16/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
2025-07-22 12:29 | 2025-07-24 13:40 | Mount Sinai Queens (MSQ) | Alyjah Luna |  | Medical Device | Ambu (via Medline) Product Recall Notification R-25-162 Ambu Inc Ambu® SPUR® Resuscitator variants with Blocked Manometer Port Recall | No |  | I have 2 cases with unaffected Lot #'s - 7/24/2025 Alyjah Luna Completed by: Alyjah Luna | Completed
|  | Beth Israel Medical Center (BIMC) |  |  | Medical Device | Volcano Corp - Visions PV .014P RX Digital IVUS Catheter, Visions PV .018 Digital IVUS Catheter - Class 2 Recall | Yes |  | Duplicate from 6/25 This recall was removed by Christopher Paradise on 8/1/2025 | Deleted
|  | Beth Israel Medical Center (BIMC) |  |  | Medical Device | Edwards Lifesciences, LLC - FEMFLEX II - Class 1 Recall | Yes |  | Received on 5/23. Completed and closed. This recall was removed by Christopher Paradise on 7/14/2025 | Deleted
|  | Mount Sinai Hospital (MSH/East) |  |  | Medical Device | Edwards Lifesciences, LLC - FEMFLEX II - Class 1 Recall | Yes |  | Received on 5/23. Completed and closed. This recall was removed by Christopher Paradise on 7/14/2025 | Deleted
|  | Mount Sinai Morningside (MSM) |  |  | Medical Device | Stryker Corporation - Strykeflow - Class 2 Recall | Yes |  | Duplicate from 7/2/25 This recall was removed by Christopher Paradise on 7/25/2025 | Deleted
2025-07-23 08:55 | 2025-07-24 06:21 | Mount Sinai-Union Square (MSUS/Downtown-Formly PACC) | Jovel Calderon Jane Bautista |  | Medical Device | Olympus Corporation of the Americas - Blue Silicone - Class 2 Recall | No |  | No products on hand. - 7/24/2025 Jovel Calderon Completed by: Jovel Calderon | Completed
2025-07-17 07:47 | 2025-07-18 09:41 | New York Eye and Ear Infirmary (NYEEI) | Jovel Calderon Jannell Maldonado |  | Medical Device | Flexicare Medical Limited - URGENT: VOLUNTARY MEDICAL DEVICE RECALL BritePro Solo Fiber Optic Laryngoscope Handles | No |  | No affected product found on site - 7/18/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
|  | Mount Sinai Hospital (MSH/East) |  |  | Medical Device | Edwards Lifesciences, LLC - Optisite Arterial Cannula - Class 1 Recall | Yes |  | Received on 5/23. Completed and closed. This recall was removed by Christopher Paradise on 7/14/2025 | Deleted
|  | New York Eye and Ear Infirmary (NYEEI) |  |  | Medical Device | Flexicare Medical (Dongguan) Ltd. - BritePro Solo Single Use Fiber Optic Laryngoscope Handle - Class 1 Recall | Yes |  | duplicate This recall was removed by Jovel Calderon on 8/4/2025 | Deleted
|  | Mount Sinai Hospital (MSH/East) |  |  | Medical Device | CooperSurgical, Inc. - Endosee System Convenience Kit with IV Tube - Class 2 Recall | Yes |  | No affected items found in POH at MSH This recall was removed by Christopher Paradise on 7/22/2025 | Deleted
|  | Mount Sinai West (MSW) |  |  | Medical Device | Siemens Healthcare Diagnostics Inc - RAPIDPoint 500 Systems Wash/Waste Cartridge - Class 2 Recall | Yes |  | No affected items found in POH This recall was removed by Christopher Paradise on 8/4/2025 | Deleted
|  | Mount Sinai West (MSW) |  |  | Medical Device | Given Imaging Ltd. - Bravo CF capsule delivery device - Class 1 Recall | Yes |  | COMPLETED ON 6/20 NO PRODUCT FOUND This recall was removed by Rebeca Melendez on 7/14/2025 | Deleted
2025-07-23 08:55 | 2025-07-26 04:30 | Mount Sinai Queens (MSQ) | Eligio Rios Iii Michael Chu |  | Medical Device | Olympus Corporation of the Americas - Blue Silicone - Class 2 Recall | No | 1 | found one in the unit and ordered a new one to replace. - 7/25/2025 Michael Chu Completed by: Christopher Paradise | Completed
2025-07-22 08:19 | 2025-07-28 10:04 | New York Eye and Ear Infirmary (NYEEI) | Minimol Shebeen |  | Medical Device | AbbVie - URGENT MEDICAL DEVICE RECALL - UPDATE Keller FunnelTM 2 | No |  | This recall was circulated back in March 2024. They have now added a lot number to the affected lots. Please check for lot number: 24C12C - 7/22/2025 Christopher Paradise We do not have any lot numbers - Janell - 7/28/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
2025-07-23 08:55 | 2025-07-23 15:01 | South Nassau Communities Hospital (MSSN) | Camille Cox |  | Medical Device | Olympus Corporation of the Americas - Blue Silicone - Class 2 Recall | No |  | No affected items - 7/23/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
2025-07-22 08:24 | 2025-07-23 13:26 | Icahn School of Medicine (ISOM) | Sharon Bankasingh |  | Medical Device | CooperSurgical, Inc. - Endosee System Convenience Kit - Class 2 Recall | Yes |  | We do not have the Endosee system kit in office. Thanks, Sharon - 7/23/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
2025-07-17 07:47 | 2025-07-24 14:24 | Mount Sinai West (MSW) | Carmelo Calderon Astrid Lopez Jason Rodriguez |  | Medical Device | Flexicare Medical Limited - URGENT: VOLUNTARY MEDICAL DEVICE RECALL BritePro Solo Fiber Optic Laryngoscope Handles | No |  | No affected lots found in stock. Carmelo. - 7/24/2025 Carmelo Calderon Completed by: Carmelo Calderon | Completed
2025-07-31 15:41 | 2025-08-01 06:34 | Mount Sinai Hospital (MSH/East) | Rebeca Melendez Christopher Paradise |  | Medical Device | Philips - UPDATED: URGENT Medical Device Correction Trilogy Evo, Trilogy Evo 02, Trilogy Evo Universal, and Trilogy EV300 Flow Sensor Nebulized Aerosol Deposition | No |  | No affected items found in POH - 8/1/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
2025-07-16 07:53 | 2025-07-22 08:56 | Mount Sinai Morningside (MSM) | Sheryl Mumford - Biomed Khadijah Awuma Sheryl Mumford |  | Medical Device | BD - URGENT MEDICAL DEVICE RECALL (CORRECTION) (MMS-25-5311-FA) Revised performance data for BD Alaris™ Pump Module Model 8100 with Guardrails™ Suite MX software when used with a subset of BD Alaris™ Compatible Pump Infusion Sets (Attachment A) Type of Action: Medical Device Correction | No |  | No affected items on site - 7/22/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
|  | Beth Israel Medical Center (BIMC) |  |  | Medical Device | CooperSurgical, Inc. - Endosee System Convenience Kit with IV Tube - Class 2 Recall | Yes |  | Site N/A This recall was removed by Christopher Paradise on 7/22/2025 | Deleted
|  | Mount Sinai Morningside (MSM) |  |  | Medical Device | Boston Scientific Corporation - Torflex Transseptal Guiding Sheath Kit - Class 2 Recall | Yes |  | Received on 6/17. Completed and closed. This recall was removed by Christopher Paradise on 7/14/2025 | Deleted
|  | Beth Israel Medical Center (BIMC) |  |  | Medical Device | Olympus Corporation of the Americas - Blue Silicone - Class 2 Recall | Yes |  | Duplicate This recall was removed by Christopher Paradise on 7/23/2025 | Deleted
|  | Mount Sinai Hospital (MSH/East) |  |  | Medical Device | B BRAUN MEDICAL INC - Extension Set - Class 1 Recall | Yes |  | DUPLICATE This recall was removed by Christopher Paradise on 8/8/2025 | Deleted
2025-07-14 09:17 | 2025-07-16 09:30 | Mount Sinai Morningside (MSM) | Mark Lydia Adelsi Soto Jazzmin Gamble Erika Saldana |  | Medical Device | Getinge - URGENT MEDICAL DEVICE – Correction - Atrium Ocean, Oasis, and Express Chest Drains | No | 50 | information forwarded - 7/16/2025 Adelsi Soto Completed by: Adelsi Soto | Completed
2025-07-16 07:53 | 2025-07-16 12:45 | Mount Sinai Doctors | Raquel Padilla Torrian Lee Russell Smith - GE Russell Smith - MS |  | Medical Device | BD - URGENT MEDICAL DEVICE RECALL (CORRECTION) (MMS-25-5311-FA) Revised performance data for BD Alaris™ Pump Module Model 8100 with Guardrails™ Suite MX software when used with a subset of BD Alaris™ Compatible Pump Infusion Sets (Attachment A) Type of Action: Medical Device Correction | No |  | not affected by this recall - 7/16/2025 Russell Smith - MS Completed by: Christopher Paradise | Completed
|  | Beth Israel Medical Center (BIMC) |  |  | Medical Device | Flexicare Medical (Dongguan) Ltd. - BritePro Solo Single Use Fiber Optic Laryngoscope Handle - Class 1 Recall | Yes |  | hospital is now closed This recall was removed by Jovel Calderon on 8/4/2025 | Deleted
|  | Mount Sinai Brooklyn Hospital (MSB) |  |  | Medical Device | Flexicare Medical (Dongguan) Ltd. - BritePro Solo SingeUse Fiber Optic Handle and Blade - Class 1 Recall | Yes |  | Duplicate from July 2025 This recall was removed by Christopher Paradise on 8/4/2025 | Deleted
|  | Mount Sinai Queens (MSQ) |  |  | Medical Device | Stryker Corporation - Stryker SmartPump Tourniquet - Class 2 Recall | Yes |  | DUPLICATE : NO AFFECTED PRODUCTS IN POH COMPLETED ON 7/3 This recall was removed by Rebeca Melendez on 8/8/2025 | Deleted
2025-07-22 12:29 | 2025-07-29 14:51 | Mount Sinai West (MSW) | Astrid Lopez |  | Medical Device | Ambu (via Medline) Product Recall Notification R-25-162 Ambu Inc Ambu® SPUR® Resuscitator variants with Blocked Manometer Port Recall | No |  | no affected product - 7/29/2025 Astrid Lopez Completed by: Astrid Lopez | Completed
2025-07-16 07:53 | 2025-07-16 09:09 | Mount Sinai Hospital (MSH/East) | Raquel Padilla Torrian Lee Russell Smith - MS |  | Medical Device | BD - URGENT MEDICAL DEVICE RECALL (CORRECTION) (MMS-25-5311-FA) Revised performance data for BD Alaris™ Pump Module Model 8100 with Guardrails™ Suite MX software when used with a subset of BD Alaris™ Compatible Pump Infusion Sets (Attachment A) Type of Action: Medical Device Correction | No |  | MSQ, MSH, and MSSN have been identified as affected sites in the BD Urgent Medical Device Recall (Correction), Reference Number MMS-25-5311-FA. All three facilities have been notified accordingly. This communication serves to ensure awareness across impacted locations and promote timely coordination in response to the recall. - 7/16/2025 Russell Smith - MS Completed by: Christopher Paradise | Completed
|  | Mount Sinai Brooklyn Hospital (MSB) |  |  | Medical Device | Ortho-Clinical Diagnostics, Inc. - VITROS Chemistry Products Performance Verifier II - Class 2 Recall | Yes |  | No affected items found in POH. Shared with Lab for awareness. This recall was removed by Christopher Paradise on 7/18/2025 | Deleted
|  | Mount Sinai Hospital (MSH/East) |  |  | Medical Device | Flexicare Medical (Dongguan) Ltd. - BritePro Solo Mini SingleUse Fiber Optic Laryngoscope Handle - Class 1 Recall | Yes |  | Duplicate from July 2025 This recall was removed by Christopher Paradise on 8/4/2025 | Deleted
2025-07-30 16:19 | 2025-08-01 12:07 | South Nassau Communities Hospital (MSSN) | Camille Cox |  | Medical Device | Stryker - Medical Device Recall - SafeAir Smoke Evacuation Pencil | No |  | All 3 depts. inventory was checked E3, OR and L&D. No affected lots were found. - 8/1/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
|  | Mount Sinai Brooklyn Hospital (MSB) |  |  | Medical Device | Olympus Corporation of the Americas - Blue Silicone - Class 2 Recall | Yes |  | Duplicate This recall was removed by Christopher Paradise on 7/23/2025 | Deleted
|  | Beth Israel Medical Center (BIMC) |  |  | Medical Device | Ortho-Clinical Diagnostics, Inc. - VITROS Chemistry Products Performance Verifier I - Class 2 Recall | Yes |  | Hospital is closed This recall was removed by Jovel Calderon on 7/18/2025 | Deleted
|  | Mount Sinai Hospital (MSH/East) |  |  | Medical Device | Stryker Corporation - Stryker SmartPump Tourniquet - Class 2 Recall | Yes |  | DUPLICATE : NO AFFECTED PRODUCTS IN POH COMPLETED ON 7/3 This recall was removed by Rebeca Melendez on 8/8/2025 | Deleted
2025-08-06 12:06 |  | Mount Sinai Hospital (MSH/East) | Maura Carpo Jason Smith Ishmael Cobbold |  | Medical Device | Cook Medical - MEDICAL DEVICE RECALL EFF,􀀁TIVENESS CHECK URGENT: PROMPT RESPO􀀁SE REQUIRED - Cook Beacon® Tip 5.0 Fr Angi1ographic Catheters | No |  |  | In Progress
2025-07-23 08:55 | 2025-07-28 14:39 | Mount Sinai West (MSW) | Carmelo Calderon Astrid Lopez Lisa Webb Elena Sasu |  | Medical Device | Olympus Corporation of the Americas - Blue Silicone - Class 2 Recall | No |  | As per Astrid Lopez no affected product in stock - 7/28/2025 Carmelo Calderon Completed by: Carmelo Calderon | Completed
2025-07-14 09:17 | 2025-07-21 10:14 | Mount Sinai Hospital (MSH/East) | Lucy Duffy Maura Carpo Jason Smith Melody Dolphy Jesus Mercado Pedro Rosario Jessica Williams Ishmael Cobbold Leandro Romero Mary Rodriguez |  | Medical Device | Getinge - URGENT MEDICAL DEVICE – Correction - Atrium Ocean, Oasis, and Express Chest Drains | No | 6 | this is not a recall….there is no action. Yes we have them, I don’t know how many. - Jessica Williams - 7/21/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
2025-07-24 06:10 | 2025-07-24 06:11 | Mount Sinai Hospital (MSH/East) | Rebeca Melendez Christopher Paradise |  | Medical Device | Intuitive - Update to the da Vinci X/Xi Arm Drape (PN 470015} | No |  | No affected product found. - 7/24/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
2025-07-16 07:53 | 2025-07-16 09:04 | South Nassau Communities Hospital (MSSN) | Glenn Havens |  | Medical Device | BD - URGENT MEDICAL DEVICE RECALL (CORRECTION) (MMS-25-5311-FA) Revised performance data for BD Alaris™ Pump Module Model 8100 with Guardrails™ Suite MX software when used with a subset of BD Alaris™ Compatible Pump Infusion Sets (Attachment A) Type of Action: Medical Device Correction | No |  | We do not use any of the infusion sets listed on this recall.- Anthony Battaglia - 7/16/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
|  | Mount Sinai Hospital (MSH/East) |  |  | Medical Device | Edwards Lifesciences, LLC - Optisite Arterial Cannula - Class 1 Recall | Yes |  | Received on 5/23. Completed and closed. This recall was removed by Christopher Paradise on 7/14/2025 | Deleted
|  | Mount Sinai Hospital (MSH/East) |  |  | Medical Device | Abiomed, Inc. - Automated Impella Controller - Class 1 Recall | Yes |  | No affected product found in POH This recall was removed by Christopher Paradise on 7/29/2025 | Deleted
|  | Mount Sinai Brooklyn Hospital (MSB) |  |  | Medical Device | Stryker Corporation - Stryker SmartPump Tourniquet - Class 2 Recall | Yes |  | DUPLICATE : NO AFFECTED PRODUCTS IN POH COMPLETED ON 7/3 This recall was removed by Rebeca Melendez on 8/8/2025 | Deleted
|  | Mount Sinai-Union Square (MSUS/Downtown-Formly PACC) |  |  | Medical Device | Olympus Corporation of the Americas - Blue Silicone - Class 2 Recall | Yes |  | Duplicate This recall was removed by Christopher Paradise on 7/23/2025 | Deleted
|  | Mount Sinai Morningside (MSM) |  |  | Medical Device | Abiomed, Inc. - Automated Impella Controller - Class 1 Recall | Yes |  | No affected product found in POH This recall was removed by Christopher Paradise on 7/29/2025 | Deleted
|  | Mount Sinai Morningside (MSM) |  |  | Medical Device | Flexicare Medical (Dongguan) Ltd. - BritePro Solo Single Use Fiber Optic Laryngoscope Handle - Class 1 Recall | Yes |  | Duplicate from July 2025 This recall was removed by Christopher Paradise on 8/4/2025 | Deleted
|  | Mount Sinai West (MSW) |  |  | Medical Device | Flexicare Medical (Dongguan) Ltd. - Britepro Solo Stubby SingleUse Fiber Optic Laryngoscope Handle - Class 1 Recall | Yes |  | Duplicate from July 2025 This recall was removed by Christopher Paradise on 8/4/2025 | Deleted
2025-07-23 08:55 | 2025-07-23 13:25 | Icahn School of Medicine (ISOM) | Juan Mercado Natalia Smith |  | Medical Device | Olympus Corporation of the Americas - Blue Silicone - Class 2 Recall | No |  | We do not have the Blue Silicone seal 66-12Fr. Model Number CS-B612.- Natalia - 7/23/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
2025-08-06 12:06 |  | Mount Sinai West (MSW) | Carmelo Calderon Astrid Lopez Lisa Webb |  | Medical Device | Cook Medical - MEDICAL DEVICE RECALL EFF,􀀁TIVENESS CHECK URGENT: PROMPT RESPO􀀁SE REQUIRED - Cook Beacon® Tip 5.0 Fr Angi1ographic Catheters | No |  |  | In Progress
2025-07-16 07:53 | 2025-07-16 12:47 | New York Eye and Ear Infirmary (NYEEI) | Russell Smith - GE Russell Smith - MS |  | Medical Device | BD - URGENT MEDICAL DEVICE RECALL (CORRECTION) (MMS-25-5311-FA) Revised performance data for BD Alaris™ Pump Module Model 8100 with Guardrails™ Suite MX software when used with a subset of BD Alaris™ Compatible Pump Infusion Sets (Attachment A) Type of Action: Medical Device Correction | No |  | NYEE is not affected by this recall - 7/16/2025 Russell Smith - MS Completed by: Christopher Paradise | Completed
|  | Mount Sinai Hospital (MSH/East) |  |  | Medical Device | Olympus Corporation of the Americas - Ellik Latex Bulb - Class 2 Recall | Yes |  | No affected product found in POH This recall was removed by Christopher Paradise on 7/23/2025 | Deleted
2025-07-23 08:48 | 2025-07-23 12:14 | Mount Sinai Morningside (MSM) | Adelsi Soto Jazzmin Gamble Erika Saldana |  | Medical Device | Olympus Corporation of the Americas - Blue Silicone - Class 2 Recall | Yes |  | Olympus Corporation of the Americas PO Date: 3/28/25 PO#: W1045973 Item #: 190A - 7/23/2025 Christopher Paradise 1 PO found with affected product in 2025 - 7/23/2025 Christopher Paradise 1 PO found with affected product in 2025 - 7/23/2025 Christopher Paradise no affected products found - 7/23/2025 Erika Saldana Completed by: Christopher Paradise | Completed
2025-07-16 07:53 | 2025-07-30 07:01 | Mount Sinai Queens (MSQ) | Franz Dunel Alexandra Zamudio |  | Medical Device | BD - URGENT MEDICAL DEVICE RECALL (CORRECTION) (MMS-25-5311-FA) Revised performance data for BD Alaris™ Pump Module Model 8100 with Guardrails™ Suite MX software when used with a subset of BD Alaris™ Compatible Pump Infusion Sets (Attachment A) Type of Action: Medical Device Correction | No | 422 | Recall is for the tube not for the actual pump - 7/21/2025 Franz Dunel As a follow up, materials was contacted and said that we did not have the affected tubing; Just to be sure the pharmacy team pulled a sample of each of the tubings used on the peds units and we did not find any of the affected tubing, so it looks like we are not affected and do not have to address the issues/concerns in the recall letter. Thanks, Simone - 7/30/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
2025-07-22 08:19 | 2025-07-22 09:49 | Mount Sinai Morningside (MSM) | Adelsi Soto Jazzmin Gamble Erika Saldana |  | Medical Device | AbbVie - URGENT MEDICAL DEVICE RECALL - UPDATE Keller FunnelTM 2 | No |  | This recall was circulated back in March 2024. They have now added a lot number to the affected lots. Please check for lot number: 24C12C - 7/22/2025 Christopher Paradise No Affected products found - 7/22/2025 Erika Saldana Completed by: Adelsi Soto | Completed
2025-07-23 12:02 | 2025-07-29 14:41 | The Blavatnik Family Chelsea Medical Center | Jane Bautista |  | Medical Device | BD URGENT Medical Device Recall SureStep™ Foley Tray Systems: Bardex® I.C. Complete Care® Infection Control Foley Catheter Tray; Lubri-Sil® I.C. Complete Care® Infection Control Foley Catheter Tray; BARD® Lubri-Sil® Foley Catheter Tray | No | 1 | After looking through our stock, we have material number 53163031 at our clinic. The items have not been used so I will place them to the side. Please let us know what needs to be done. - 7/29/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
|  | Mount Sinai West (MSW) |  |  | Medical Device | Volcano Corp - Visions PV .014P RX Digital IVUS Catheter, Visions PV .018 Digital IVUS Catheter - Class 2 Recall | Yes |  | Duplicate from 6/25 This recall was removed by Christopher Paradise on 8/1/2025 | Deleted
|  | Mount Sinai Morningside (MSM) |  |  | Medical Device | Ortho-Clinical Diagnostics, Inc. - VITROS Chemistry Products Performance Verifier I - Class 2 Recall | Yes |  | No affected items found in POH. Shared with Lab for awareness. This recall was removed by Christopher Paradise on 7/18/2025 | Deleted
2025-07-24 06:10 |  | Mount Sinai Doctors |  |  | Medical Device | Intuitive - Update to the da Vinci X/Xi Arm Drape (PN 470015} | No |  | No affected product found This recall was removed by Christopher Paradise on 7/24/2025 | Deleted
|  | Beth Israel Medical Center (BIMC) |  |  | Medical Device | Flexicare Medical (Dongguan) Ltd. - BritePro Solo SingeUse Fiber Optic Handle and Blade - Class 1 Recall | Yes |  | hospital is closed no inventory This recall was removed by Jovel Calderon on 8/4/2025 | Deleted
|  | South Nassau Communities Hospital (MSSN) |  |  | Medical Device | Olympus Corporation of the Americas - Blue Silicone - Class 2 Recall | Yes |  | Duplicate This recall was removed by Christopher Paradise on 7/23/2025 | Deleted
2025-07-23 12:02 | 2025-07-25 08:53 | Mount Sinai Queens (MSQ) | Vanessa Palma-Tejeda Renato Pasion Claire Jazmin-Aquino Victor Mena Eleana Candelier Rayduan Saban |  | Medical Device | BD URGENT Medical Device Recall SureStep™ Foley Tray Systems: Bardex® I.C. Complete Care® Infection Control Foley Catheter Tray; Lubri-Sil® I.C. Complete Care® Infection Control Foley Catheter Tray; BARD® Lubri-Sil® Foley Catheter Tray | No |  | We do not have any of the affected lots of these products. - 7/25/2025 Vanessa Palma-Tejeda Completed by: Vanessa Palma-Tejeda | Completed
2025-08-06 10:14 |  | Mount Sinai Hospital (MSH/East) | Juanita Jordan Arbinda Gurung |  | Medical Device | Boston Scientific - Early Alert: Defibrillation Lead Issue from BostonScientific Endotak Reliance | No |  | I found 1 Purchase Order from 2023 with affected product. Please let me know if we have any affected product on site in Mount Sinai Hospital. - 8/6/2025 Christopher Paradise | In Progress
2025-07-22 08:19 | 2025-07-22 11:45 | Icahn School of Medicine (ISOM) | Andrew Lebron Francisco Fernandez |  | Medical Device | AbbVie - URGENT MEDICAL DEVICE RECALL - UPDATE Keller FunnelTM 2 | No |  | This recall was circulated back in March 2024. They have now added a lot number to the affected lots. Please check for lot number: 24C12C - 7/22/2025 Christopher Paradise none on hand - 7/22/2025 Francisco Fernandez Completed by: Francisco Fernandez | Completed
2025-07-16 07:53 | 2025-07-16 12:45 | Mount Sinai West (MSW) | Ernesto Dilizia Karina Rodriguez |  | Medical Device | BD - URGENT MEDICAL DEVICE RECALL (CORRECTION) (MMS-25-5311-FA) Revised performance data for BD Alaris™ Pump Module Model 8100 with Guardrails™ Suite MX software when used with a subset of BD Alaris™ Compatible Pump Infusion Sets (Attachment A) Type of Action: Medical Device Correction | No |  | West does not have Alaris pumps - 7/16/2025 Ernesto Dilizia Completed by: Christopher Paradise | Completed
|  | Mount Sinai Hospital (MSH/East) |  |  | Medical Device | Beckman Coulter, Inc. - Access Vitamin B12 Reagent - Class 2 Recall | Yes |  | No affected items found in POH This recall was removed by Christopher Paradise on 7/14/2025 | Deleted
2025-07-16 07:53 | 2025-07-16 12:46 | Mount Sinai-Union Square (MSUS/Downtown-Formly PACC) | Russell Smith - GE Russell Smith - MS |  | Medical Device | BD - URGENT MEDICAL DEVICE RECALL (CORRECTION) (MMS-25-5311-FA) Revised performance data for BD Alaris™ Pump Module Model 8100 with Guardrails™ Suite MX software when used with a subset of BD Alaris™ Compatible Pump Infusion Sets (Attachment A) Type of Action: Medical Device Correction | No |  | Mount Sinai-Union Square (MSUS/Downtown-Formly PACC) is not affected by this recall - 7/16/2025 Russell Smith - MS Completed by: Christopher Paradise | Completed
|  | Mount Sinai Morningside (MSM) |  |  | Medical Device | Olympus Corporation of the Americas - Glass Body For 194 Ellik Evacuator - Class 2 Recall | Yes |  | No affected product found on POH This recall was removed by Christopher Paradise on 7/23/2025 | Deleted
|  | South Nassau Communities Hospital (MSSN) |  |  | Medical Device | Stryker Corporation - Stryker SmartPump Tourniquet - Class 2 Recall | Yes |  | DUPLICATE : NO AFFECTED PRODUCTS IN POH COMPLETED ON 7/3 This recall was removed by Rebeca Melendez on 8/8/2025 | Deleted
2025-07-14 09:17 | 2025-07-21 08:27 | Mount Sinai Brooklyn Hospital (MSB) | Karen Grant Manuel Pimentel Venetia Walker-Cowan Belinda Baril |  | Medical Device | Getinge - URGENT MEDICAL DEVICE – Correction - Atrium Ocean, Oasis, and Express Chest Drains | No |  | No affected product at MSB - 7/21/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
2025-07-16 13:26 | 2025-07-21 10:48 | Mount Sinai Hospital (MSH/East) | Maura Carpo Jason Smith Ishmael Cobbold |  | Medical Device | GETINGE - Vasoview HemoPro 2 Endoscopic Vessel Harvesting (EVH) System - CORRECTION UPDATE | No | 4 | Updated correction letter was sent to OR Education and Chief of Perfusion to distribute to any affected parties. - 7/21/2025 Jason Smith Completed by: Jason Smith | Completed
|  | Mount Sinai Hospital (MSH/East) |  |  | Medical Device | Olympus Corporation of the Americas - Blue Silicone - Class 2 Recall | Yes |  | Duplicate This recall was removed by Christopher Paradise on 7/23/2025 | Deleted
|  | Mount Sinai Morningside (MSM) |  |  | Medical Device | Boston Scientific Corporation - VersaCross Access Solution - Class 2 Recall | Yes |  | Received on 6/17. Completed and closed with no affected items found This recall was removed by Christopher Paradise on 7/14/2025 | Deleted
2025-07-29 06:31 | 2025-07-30 07:37 | Mount Sinai Hospital (MSH/East) | Jesus Mercado Pedro Rosario Leandro Romero |  | Medical Device | Vantive - Urgent Medical Device Correction - PrisMax & Thermax Alarms | No |  | Sent email to Erin Figuroa who overseas Women & Children to see if she has contacts for the affected departments. - 7/30/2025 Christopher Paradise The Prismax devices (3) were purchased for our Pediatric ICU (1184 5th Ave, 3rd floor). Location of the machines. We too received the Urgent Medical Device Correction Notification alert from Elizabeth Ziefer who is included in the email. From what we have reviewed with our PICU team is that there is nothing to be done at this time but to monitor the alerts and troubleshoot as we would normally do if the alert popped up on the machine. Please advise if there is anything in addition that is required to monitor the Prismax devices or any updated that need to be completed for them to alleviate the alerts. I have also included Juliana the NM for the PICU. Thank you Svetlana Streltsova-Verma, DNP, RN, CCRN, NE-BC - 7/30/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
2025-07-22 08:30 | 2025-07-23 13:26 | Icahn School of Medicine (ISOM) | Sharon Bankasingh |  | Medical Device | CooperSurgical, Inc. - Endosee System Convenience Kit with IV Tube - Class 2 Recall | Yes |  | We do not have the Endosee system kit in office. Thanks, Sharon - 7/23/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
|  | Mount Sinai Morningside (MSM) |  |  | Medical Device | Edwards Lifesciences, LLC - Optisite Arterial Cannula - Class 1 Recall | Yes |  | Received on 5/23. Completed and closed. This recall was removed by Christopher Paradise on 7/14/2025 | Deleted
2025-07-22 08:19 | 2025-07-24 14:29 | Mount Sinai West (MSW) | Carmelo Calderon Astrid Lopez Lisa Webb |  | Medical Device | AbbVie - URGENT MEDICAL DEVICE RECALL - UPDATE Keller FunnelTM 2 | No |  | This recall was circulated back in March 2024. They have now added a lot number to the affected lots. Please check for lot number: 24C12C - 7/22/2025 Christopher Paradise As per Astrid Lopez no effected product in stock. - 7/24/2025 Carmelo Calderon Completed by: Carmelo Calderon | Completed
2025-07-22 12:29 | 2025-07-29 06:55 | Mount Sinai Hospital (MSH/East) | Jesus Mercado Pedro Rosario Gary Oldenberg Milton Mancero Leandro Romero |  | Medical Device | Ambu (via Medline) Product Recall Notification R-25-162 Ambu Inc Ambu® SPUR® Resuscitator variants with Blocked Manometer Port Recall | No |  | No affected product found on site - 7/29/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
|  | Beth Israel Medical Center (BIMC) |  |  | Medical Device | Stryker Corporation - Stryker SmartPump Tourniquet - Class 2 Recall | Yes |  | Hospital is now closed This recall was removed by Jovel Calderon on 8/8/2025 | Deleted
2025-07-17 07:47 | 2025-07-22 08:04 | Mount Sinai Morningside (MSM) | Adelsi Soto Jazzmin Gamble Tricia Carrington Erika Saldana |  | Medical Device | Flexicare Medical Limited - URGENT: VOLUNTARY MEDICAL DEVICE RECALL BritePro Solo Fiber Optic Laryngoscope Handles | No |  | no affected lots found - 7/17/2025 Erika Saldana Completed by: Christopher Paradise | Completed
|  | Mount Sinai Morningside (MSM) |  |  | Medical Device | Flexicare Medical (Dongguan) Ltd. - BritePro Solo SingeUse Fiber Optic Handle and Blade - Class 1 Recall | Yes |  | Duplicate from July 2025 This recall was removed by Christopher Paradise on 8/4/2025 | Deleted
2025-07-11 12:53 | 2025-07-11 12:54 | Mount Sinai Hospital (MSH/East) | Rebeca Melendez Christopher Paradise |  | Medical Device | Masimo Recall for Radical-7 Pulse CO-Oximeter batteries | No |  | Nothing found in POH. Sent to Biomed for awareness. - 7/11/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
2025-07-16 07:53 | 2025-07-16 09:05 | Mount Sinai Brooklyn Hospital (MSB) | Russell Smith - GE Ronald Narchet Russell Smith - MS |  | Medical Device | BD - URGENT MEDICAL DEVICE RECALL (CORRECTION) (MMS-25-5311-FA) Revised performance data for BD Alaris™ Pump Module Model 8100 with Guardrails™ Suite MX software when used with a subset of BD Alaris™ Compatible Pump Infusion Sets (Attachment A) Type of Action: Medical Device Correction | No |  | MSB is not affected by this recall - 7/16/2025 Russell Smith - MS Completed by: Christopher Paradise | Completed
|  | Beth Israel Medical Center (BIMC) |  |  | Medical Device | Siemens Healthcare Diagnostics Inc - RAPIDPoint 500 Systems Wash/Waste Cartridge - Class 2 Recall | Yes |  | No affected items found in POH This recall was removed by Christopher Paradise on 8/4/2025 | Deleted
2025-07-22 08:19 | 2025-07-23 15:01 | South Nassau Communities Hospital (MSSN) | Camille Cox |  | Medical Device | AbbVie - URGENT MEDICAL DEVICE RECALL - UPDATE Keller FunnelTM 2 | No |  | This recall was circulated back in March 2024. They have now added a lot number to the affected lots. Please check for lot number: 24C12C - 7/22/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
2025-08-04 05:59 |  | Mount Sinai Hospital (MSH/East) |  |  | Medical Device | Remel, Inc - Blood Agar,5 Sheep Blood - Class 2 Recall | Yes |  | No affected items found in POH This recall was removed by Christopher Paradise on 8/4/2025 | Deleted
|  | Mount Sinai Hospital (MSH/East) |  |  | Medical Device | Ortho-Clinical Diagnostics, Inc. - VITROS Chemistry Products Performance Verifier II - Class 2 Recall | Yes |  | No affected items found in POH. Shared with Lab for awareness. This recall was removed by Christopher Paradise on 7/18/2025 | Deleted
|  | Mount Sinai Hospital (MSH/East) |  |  | Medical Device | Ortho-Clinical Diagnostics, Inc. - VITROS Chemistry Products Performance Verifier I - Class 2 Recall | Yes |  | No affected items found in POH. Shared with Lab for awareness. This recall was removed by Christopher Paradise on 7/18/2025 | Deleted
2025-08-06 12:16 | 2025-08-06 12:17 | Mount Sinai Hospital (MSH/East) | Rebeca Melendez Christopher Paradise |  | Medical Device | MEDLINE RENEWAL Electrophysiology Catheter Issue from Medline ReNewal | No |  | NO AFFECTED PRODUCTS FOUND IN POH - 8/6/2025 Rebeca Melendez Completed by: Rebeca Melendez | Completed
|  | Mount Sinai Hospital (MSH/East) |  |  | Medical Device | CareFusion 303, Inc. - BD Pyxis Products - Class 2 Recall | Yes |  | Sent to Pharmacy This recall was removed by Christopher Paradise on 7/18/2025 | Deleted
|  | Mount Sinai West (MSW) |  |  | Medical Device | Olympus Corporation of the Americas - Blue Silicone - Class 2 Recall | Yes |  | Duplicate This recall was removed by Christopher Paradise on 7/23/2025 | Deleted
2025-07-28 12:40 | 2025-07-28 13:17 | Mount Sinai Hospital (MSH/East) | Rebeca Melendez Christopher Paradise |  | Medical Device | Philips - URGENT MEDICAL DEVICE CORRECTION - Ultrasound Transducers (Accessory to Philips Ultrasound System) Expected Useful Life and Date of Manufacturing Labeling | No | 1 | I am not finding any of the affected transducer products in our POH. I will save the record. - 7/28/2025 Christopher Paradise We do have some, we knew this was coming. We are working with GE to swap out probes. Just wanted you to have the recall since they finally sent it out. Jessica Williams - 7/28/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
|  | Mount Sinai-Union Square (MSUS/Downtown-Formly PACC) |  |  | Medical Device | Flexicare Medical (Dongguan) Ltd. - BritePro Solo SingeUse Fiber Optic Handle and Blade - Class 1 Recall | Yes |  | no affected lot# on inventory This recall was removed by Jovel Calderon on 8/4/2025 | Deleted
2025-07-22 08:19 | 2025-07-28 10:05 | The Blavatnik Family Chelsea Medical Center | Jane Bautista |  | Medical Device | AbbVie - URGENT MEDICAL DEVICE RECALL - UPDATE Keller FunnelTM 2 | No |  | This recall was circulated back in March 2024. They have now added a lot number to the affected lots. Please check for lot number: 24C12C - 7/22/2025 Christopher Paradise The recall notification was filled up and was return to the vendor already, see attached. - Jovel - 7/28/2025 Christopher Paradise Completed by: Christopher Paradise | Completed
|  | Icahn School of Medicine (ISOM) |  |  | Medical Device | Olympus Corporation of the Americas - Blue Silicone - Class 2 Recall | Yes |  | Duplicate This recall was removed by Christopher Paradise on 7/23/2025 | Deleted
2025-08-08 06:47 |  | Mount Sinai Hospital (MSH/East) | Juanita Jordan Arbinda Gurung |  | Medical Device | Philips - Urgent Medical Device CORRECTION - Philips Azurion and Allura with 1 Phase UPS Potential Loss of System Functionality | No |  |  | In Progress
|  | Mount Sinai Hospital (MSH/East) |  |  | Medical Device | Edwards Lifesciences, LLC - FEMFLEX II - Class 1 Recall | Yes |  | Received on 5/23. Completed and closed. This recall was removed by Christopher Paradise on 7/14/2025 | Deleted
|  | Mount Sinai Hospital (MSH/East) |  |  | Medical Device | Flexicare Medical (Dongguan) Ltd. - BritePro Solo Single Use Fiber Optic Laryngoscope Handle - Class 1 Recall | Yes |  | Duplicate from July 2025 This recall was removed by Christopher Paradise on 8/4/2025 | Deleted
2025-07-14 09:17 | 2025-07-21 09:09 | Mount Sinai West (MSW) | James Fox Carmelo Calderon Astrid Lopez Kenrick McLeod Lisa Webb |  | Medical Device | Getinge - URGENT MEDICAL DEVICE – Correction - Atrium Ocean, Oasis, and Express Chest Drains | No |  | NONE IN STOCK - 7/21/2025 Kenrick McLeod Completed by: Carmelo Calderon | Completed
CONFIDENTIAL: This information is intended for use by MD Buyline members only.